INTRODUCTION

DAQ guide ecogas setup

This guide is meant to explain how to use the ecogas data taking setup and it is divided in the following sections: 1) hardware setup, which contains information on the hardware used in the setup 2) software setup, which describes briefly how the code for this setup works 3) type of runs, which describes the possible runs that cane be started 4) analysis, describes the analysis code and 5) possible errors gives a list of the most common errors that can appear during data taking.

Setup developed by Luca Quaglia and Sara Garetti in 2023 at the INFN Turin laboratory.

Guide is still in evolution since the setup is still evolving.

HARDWARE SETUP

High Voltage (HV) module

A CAEN SY1527 HV mainframe is used to supply the HV to the detectors. A series of A1526P (for the RPCs) and A1733N (for the photomultipliers) are installed in the back of the mainframe. The IP address of the mainframe is 90.147.203.174 and the telnet protocol can be used to connect to it from any PC connected to the INFN network (by inserting in terminal the following command: telnet 90.147.203.174 1527).

The PMTs used as trigger are named PMT_GAS_1...4 and are connected to slot 11, channels 0...4 while the RPC currently under test is named ECOgas and is connected to slot 7 channel 1.

A Python wrapper of the CAEN HV library is used to communicate with the module (will be described later on in the guide). The files needed for this to work are the following:
1) /home/pcald32/labStrada/DAQ/currentScan/CAEN.py (python class to wrap C++ functions)
2) /usr/lib64/libcaenhvwrapper.so.6.3 (CAEN library, which contains the C++ functions)

Gas mixing system

The gas mixing system is located in the gas room outside of the lab. The gas room contains all the bottles and the mixing system. The bottles are equipped with two pressure regulators, one attached directly to the bottle and the other one just before the mass flow controller.

The pressure after the bottle is already set and should not be touched, for what concerns the pressure before the mass flow controllers, it should be set to ~ 1 bar for each gas (by rotating the valve towards the “+” sign).

A total of 4 Bronkhorst mass flow controllers (??) is used to create the desired gas mixture. For the moment, channel 1 is connected to Ar, channel 2 to R134a, channel 3 to i-C4H10 and channel 4 to SF6. The power is provided through a power converter attached on the metal support frame (Fig. xx). The MFCs are connected to on other through ethernet cables while the connection to the laptop is executed using an RS232 cable (between MFC1 and the pc).

The MFCs are operated using the flowDDE and flowView tools, already installed on the PC. The procedure to operate the MFCs is the following:
1) Switch on the pc and connect the RS232 cable through the RS232-USB adapter
2) Open the flowDDE software >> Communication >> Open communication (F3) to start the communication between the pc and the MFC and wait for the display to show “flowDDE is active and ready for any client”
3) Open as many flowView panels as the MFC and set the flow in each of them. Note that you need to wait for a panel to be fully opened before opening a new one. The flow can be set either in terms of setpoint (in ml/min) or as a percentage of the maximum flow in each MFC. Moreover the MFCs can be arranged in a master/slave configuration. In order to do that, in the flowview panel >> advanced >> Setpoint/Control Mode >> 0. Bus RS232 for the master while 2. FB/RS232 slave for the slaves. Moreover, in the master flowView panel there is the “node”, this number has to be put also in the node box of all the slaves, otherwise the communication between master and slaves cannot work. The last parameter to set is

A python script “….py” is available at “….” and it is used to cycle on the four MFCs and measure both the gas flow as well as the setpoint (both in ml/min). This data is then sent to the mysql database that is running on the DAQ machine in the lab. This is used to monitor the gas flow and to issue a software interlock to the HV and stop the ongoing run. Actually the interlock is issued by another pyhton script (gasKill.py) that is running on the main DAQ PC. If the gas flow goes above/below the setpoint by 10% for more than 60 seconds the HV is killed and the ongoing measurement is stopped. In this case, a Telegram bot sends a message on the “labStradaUpdates” channel (https://t.me/labStradaUpdates). This bot also sends messages when an error is observed in Grafana (i.e. if the gas flow goes below some limits that are set in Grafana as well). Still to be implemented on the DAQ PC: before issuing the interlock the ongoing process should be stopped (find PID and kill the process)

For the moment we have an issue with the WiFi connection of the gas pc since the INFN-web network is not stable, the pc randomly disconnects from the network and it doesn’t send data to the db anymore. We are trying to get an ethernet cable pulled all the way into the gas room (not likely to happen) but for now we are developing a script that checks if the pc is connected to the network (...py), if not, it should automatically connect to the network (...py) and it automatically opens the login page and logs in with the user credentials. For now the script is almost ready but we are having some issues with the SSL/TLS certificates (search for a fix ongoing). Ideally we should be able to have ethernet connection in the gas room.

RPC setup

The setup can host up to two RPC detectors, which are inserted inside of a Farady cage. The detectors should be flushed with gas in series

Trigger logic

The trigger is built using NIM modules. A total of four scintillators + PMT is used, two of them (1 and 2) are placed on top of the setup and two more (3 and 4) are placed under.

DAQ logic

SOFTWARE SETUP

TYPE OF RUNS

ANALYSIS

DATA TAKING PROCEDURE

- Check that Grafana server is running:
1) Open Firefox browser and go to the following link: localhost:3000/
2) If Grafana panel opens up to this page then the server is working. From here click the three lines next to “Home” >> Dashboards >> labStrada

3) If the main Grafana page is not shown in localhost then it means that the Grafana server is not running. To start it you should open a terminal and type (password is pcald32):

- sudo systemctl daemon-reload → to reload the daemon
- sudo systemctl start grafana-server → to restart the server
- sudo systemctl status grafana-server → to check the status (output of this command should no be empty)

4) Once this is done you can try to connect to Grafana again and this should work

- Check that MySQL server is running:
1) Open Firefox browser and connect to the following link: http://localhost/phpmyadmin/
2) If it opens to the php myadmin main page it means the server is working and you can proceed with the data taking

3) If it is not working you need to restart the server by opening a terminal and type (password is pcald32):

- sudo /etc/init.d/mysql start → to restart the server
- sudo systemctl status mysql → to check the status (output of this command should no be empty)

4) Once this is done you can try to connect to MySQL again and it should work

- Check that the environmental parameter monitor script (monitoring.py) is running, this can be checked even in Grafana (check that a point is added every 10 minutes). If the script is not running then it can be started in the following way
1) First of all you need to load the code on the Arduino, to do so open a terminal and type:

cd arduino-ide_2.0.4_Linux_64bit/ → to open the folder where the Arduino IDE (text editor to also uplad files to the arduino) is located then type:
./arduino-ide → to launch the editor

2) Open the file: /home/pcald32/labStrada/monitoring/arduino/tempPressHumiReader/tempPressHumiReader.ino

3) Make sure that only a single Arduino IDE is open, otherwise communication with the board cannot be established.
4) Make sure that esp32 dev module is selected, to do so click on Tools > Board > esp32 > ESP32 Dev Module
5) Make sure that the proper COM port is selected, to do so click on Tools > Port > /dev/ttyUSB0
6) Click on the upload button and wait for the code to be uploaded

7) After loading the code on the arudino open a terminal anc type:
cd /home/pcald32/labStrada/monitoring/python
8) python3 monitoring.py to launch the script. After 10 minutes a line such as the following should appear (also a point in the grafana environmental monitoring panel)

- Check that the gas flow monitoring script is running on the gas pc

The gas pc is located in the gas room and it can be reached using teamviewr on the DAQ machine (pcald32). For the moment the pc is connected to the wifi network INFN-web and the connection is sometimes unstable. For this reason, the pc might be disconnected and not reachable by teamviewer. If this is the case, you should go to the gas room and connect it to the nettwork (see posible errors section).

If the script is not running (i.e. no terminal is open or the script is for some reason stuck), it is possible to start it in the following way:
1) Open a terminal (start → cmd)
2) cd Documents (to change to the folder which contains the scripts)
3) python readGas.py (to start the script, in case of errors refer to the possible error section)

- Check that the HV monitoring script is running

1) If the script is running you should a point every 30 seconds on the Grafana panel monitoring the HV. If not, it means that the script is not running. To restart it open a terminal and type:

1) cd labStrada/monitoring/python (change to the folder containing the scripts)
2) python3 hvMonitor.py (to start the script, in case of errors check the possible errors section)

POSSIBLE ERRORS

- Problem: gas pc is not connected to the wifi
- Solution:
1) Go to the gas room
2) Connect the pc to the INFN web network
3) Open Internet explorer browser and go to the 1.1.1.1 website (this will open the network login page)
4) Click on “Prosegui sul sito wen (scelta non consigliata)”, this is because the certificates on the network are expire, but don’t worry, it’s safe to connect.
5) Enter credentials: username is luca.quaglia and password is Luc4*Qgl

- Problem: gas pc is not able to upload data to the db even if it is connected to the network. Running the script readGas.py produces the following error:
mettere errore qui
- Solution: this error means that the IP address of the gas pc has changed with respect to the latest address saved on the db. We created a user in the mysql db on the daq pc that is able to send data to it from another pc. In order to do so, we have saved in the credentials the IP address of the gas pc. Since this can change it can, sometimes, happen that you need to update it also on the daq pc.
1) Go on the gas pc, open a terminal and type ipconfig (the output of this command shows the current IP address)

2) Go on the daq machine to the mysql db page. Open Firefox and type http://localhost/phpmyadmin/ to open the db
3) Open the mysql db → SQL and type the following command to update the registered IP address

4) Still in the mysql db → SQL, type the following command to grant all the rights to the pcGas user
5) Try to re-launch the script on the gas pc to check if everything is working

- Problem: any script connecting to the high voltage module cannot establish the connection
- Solution: this might mean that the connection with the module is down: open a terminal on the daq pc anf type the following command: ping 90.147.203.174 → if an error message containing the words “host not reachable” or similar, it means that either the module is not powered on or the ethernet cable is loose/disconnected. Check all the physical connections/power on the mainframe and act accordingly (replug ethernet cable/power cycle the module). Wait 5 minutes for it to reboot and then try to ping the module again, the problem should now be fixed