INTRODUCTION

DAQ guide ecogas setup

This guide is meant to explain how to use the ecogas data taking setup and it is divided in the following sections: 1) hardware setup, which contains information on the hardware used in the setup 2) software setup, which describes briefly how the code for this setup works 3) type of runs, which describes the possible runs that cane be started 4) analysis, describes the analysis code and 5) possible errors gives a list of the most common errors that can appear during data taking.

Setup developed by Luca Quaglia and Sara Garetti in 2023 at the INFN Turin laboratory.

Guide is still in evolution since the setup is still evolving.

HARDWARE SETUP

High Voltage (HV) module

A CAEN SY1527 HV mainframe is used to supply the HV to the detectors. A series of A1526P (for the RPCs) and A1733N (for the photomultipliers) are installed in the back of the mainframe. The IP address of the mainframe is 90.147.203.174 and the telnet protocol can be used to connect to it from any PC connected to the INFN network (by inserting in terminal the following command: telnet 90.147.203.174 1527).

The PMTs used as trigger are named PMT_GAS_1...4 and are connected to slot 11, channels 0...4 while the RPC currently under test is named ECOgas and is connected to slot 7 channel 1.

A Python wrapper of the CAEN HV library is used to communicate with the module (will be described later on in the guide). The files needed for this to work are the following:
1) /home/pcald32/labStrada/DAQ/currentScan/CAEN.py (python class to wrap C++ functions)
2) /usr/lib64/libcaenhvwrapper.so.6.3 (CAEN library, which contains the C++ functions)

Gas mixing system

The gas mixing system is located in the gas room outside of the lab. The gas room contains all the bottles and the mixing system. The bottles are equipped with two pressure regulators, one attached directly to the bottle and the other one just before the mass flow controller.

The pressure after the bottle is already set and should not be touched, for what concerns the pressure before the mass flow controllers, it should be set to ~ 1 bar for each gas (by rotating the valve towards the “+” sign).

A total of 4 Bronkhorst mass flow controllers (??) is used to create the desired gas mixture. For the moment, channel 1 is connected to Ar, channel 2 to R134a, channel 3 to i-C4H10 and channel 4 to SF6.

RPC setup

The setup can host up to two RPC detectors, which are inserted inside of a Farady cage. The detectors should be flushed with gas in series

Trigger logic

The trigger is built using NIM modules. A total of four scintillators + PMT is used, two of them (1 and 2) are placed on top of the setup and two more (3 and 4) are placed under.

DAQ logic

SOFTWARE SETUP

TYPE OF RUNS

ANALYSIS

POSSIBLE ERRORS